## Summary
CardioWatch 287-2B — Boot Self-Test Matrix
Generated: 2026-02-07 12:51 ¦ Platform: nRF52840 Cortex-M4 ¦ Document Rev 1.0
Project | MMT287-2B_MultiColor (CardioWatch 287-2B)
MCU | Nordic nRF52840 — ARM Cortex-M4F @ 64 MHz
Flash | 1 MB Internal + 32 MB External (MT25QU256)
RAM | 256 KB
BLE Stack | S140 SoftDevice v7.x
Boot Time | ~6.05 s (to sleep mode)
Watchdog | Active during boot, fed every 100 ms
Test Source Files | mmt_selftest.c, mmt_test_ALU_FPU.c, mmt_fds_simple.c, main.c
Test Categories Overview
# | Category | Tests | Severity | Source | Description
1 | Boot Sequence | 16 | CRITICAL | main.c | Validates ordered presence of all initialization markers in boot log
2 | FDS Integrity | 3 | CRITICAL | mmt_fds_simple.c | Write/read cycle test on Flash Data Storage (5 iterations, pattern matching)
3 | Self-Test Bitmask | 9 | CRITICAL/MAJOR | mmt_selftest.c | Hardware self-test: temperature sensors, PPG, accelerometer, timing
4 | ALU/FPU Benchmark | 7 | MINOR | mmt_test_ALU_FPU.c | Processor arithmetic performance validation (int32, float, double)
5 | Boot Timing | 2 | MAJOR | main.c + mmt_tasks.c | Initialization phase and sleep mode transition timing
6 | Temperature Init | 2 | MAJOR | main.c | Temperature sensor readings at boot within valid range
7 | Hardware Init | 1 | MINOR | main.c | TMP117 I2C sensor detection status
8 | Firmware Info | 3 | INFO | main.c + FW_info.c | FW version, serial number, and reset reason logging
Total tests: 43

## Test Matrix
Test ID | Category | Test Name | Severity | Precondition | Test Method | Pass Criteria | Golden Value / Range | Source File | Source Line | Status | Actual Value | Comments
BT-SEQ-001 | Boot Sequence | Boot marker 'BOOT_START' present and in order | CRITICAL | Device powered on or reset | Parse RTT/UART log for marker pattern | Pattern "MMT application started at 0x40000" found in log, after previous markers | Pattern: "MMT application started at 0x40000" | main.c
BT-SEQ-002 | Boot Sequence | Boot marker 'INIT_PHASE_1' present and in order | CRITICAL | Device powered on or reset | Parse RTT/UART log for marker pattern | Pattern "init 1." found in log, after previous markers | Pattern: "init 1." | main.c
BT-SEQ-003 | Boot Sequence | Boot marker 'INIT_PHASE_2' present and in order | CRITICAL | Device powered on or reset | Parse RTT/UART log for marker pattern | Pattern "init 2." found in log, after previous markers | Pattern: "init 2." | main.c
BT-SEQ-004 | Boot Sequence | Boot marker 'INIT_PHASE_3' present and in order | CRITICAL | Device powered on or reset | Parse RTT/UART log for marker pattern | Pattern "init 3." found in log, after previous markers | Pattern: "init 3." | main.c
BT-SEQ-005 | Boot Sequence | Boot marker 'INIT_PHASE_4' present and in order | CRITICAL | Device powered on or reset | Parse RTT/UART log for marker pattern | Pattern "init 4." found in log, after previous markers | Pattern: "init 4." | main.c
BT-SEQ-006 | Boot Sequence | Boot marker 'INIT_PHASE_41' present and in order | CRITICAL | Device powered on or reset | Parse RTT/UART log for marker pattern | Pattern "init 41." found in log, after previous markers | Pattern: "init 41." | main.c
BT-SEQ-007 | Boot Sequence | Boot marker 'INIT_PHASE_42' present and in order | CRITICAL | Device powered on or reset | Parse RTT/UART log for marker pattern | Pattern "init 42." found in log, after previous markers | Pattern: "init 42." | main.c
BT-SEQ-008 | Boot Sequence | Boot marker 'INIT_PHASE_43' present and in order | CRITICAL | Device powered on or reset | Parse RTT/UART log for marker pattern | Pattern "init 43." found in log, after previous markers | Pattern: "init 43." | main.c
BT-SEQ-009 | Boot Sequence | Boot marker 'INIT_PHASE_44' present and in order | CRITICAL | Device powered on or reset | Parse RTT/UART log for marker pattern | Pattern "init 44." found in log, after previous markers | Pattern: "init 44." | main.c
BT-SEQ-010 | Boot Sequence | Boot marker 'FDS_TEST_START' present and in order | CRITICAL | Device powered on or reset | Parse RTT/UART log for marker pattern | Pattern "#### TEsting FDS: mmt_fds_test" found in log, after previous markers | Pattern: "#### TEsting FDS: mmt_fds_test" | main.c
BT-SEQ-011 | Boot Sequence | Boot marker 'FDS_TEST_RESULT' present and in order | CRITICAL | Device powered on or reset | Parse RTT/UART log for marker pattern | Pattern "FDS Test (OK¦NOK)" found in log, after previous markers | Pattern: "FDS Test (OK¦NOK)" | main.c
BT-SEQ-012 | Boot Sequence | Boot marker 'DFU_WAIT' present and in order | CRITICAL | Device powered on or reset | Parse RTT/UART log for marker pattern | Pattern "waiting for potential DFU" found in log, after previous markers | Pattern: "waiting for potential DFU" | main.c
BT-SEQ-013 | Boot Sequence | Boot marker 'INIT_PHASE_6' present and in order | CRITICAL | Device powered on or reset | Parse RTT/UART log for marker pattern | Pattern "init 6." found in log, after previous markers | Pattern: "init 6." | main.c
BT-SEQ-014 | Boot Sequence | Boot marker 'APP_RUNNING' present and in order | CRITICAL | Device powered on or reset | Parse RTT/UART log for marker pattern | Pattern "Running the app" found in log, after previous markers | Pattern: "Running the app" | main.c
BT-SEQ-015 | Boot Sequence | Boot marker 'FW_VERSION' present and in order | CRITICAL | Device powered on or reset | Parse RTT/UART log for marker pattern | Pattern "MMT 287-B2 STARTED FW_V" found in log, after previous markers | Pattern: "MMT 287-B2 STARTED FW_V" | main.c
BT-SEQ-016 | Boot Sequence | Boot marker 'MAIN_LOOP' present and in order | CRITICAL | Device powered on or reset | Parse RTT/UART log for marker pattern | Pattern "MAIN LOOP STARTED" found in log, after previous markers | Pattern: "MAIN LOOP STARTED" | main.c
BT-FDS-017 | FDS Integrity | FDS test execution at boot | CRITICAL | FDS initialized, QSPI flash ready | Check for FDS test marker in boot log | "TEsting FDS: mmt_fds_test" appears in log | Marker present | main.c | 502
BT-FDS-018 | FDS Integrity | FDS write/read cycle (5 iterations) | CRITICAL | FDS test started | Write test_value (0-4) + string "MMT_FDS", read back, compare | All 5 iterations: d_read.i == d_write.i AND rc == NRF_SUCCESS | Values 0-4, string "MMT_FDS", file_id=0x8123, key=0x7234 | mmt_fds_simple.c | 700-720
BT-FDS-019 | FDS Integrity | FDS overall result | CRITICAL | FDS test completed | Check for "FDS Test OK" or "FDS Test NOK" in log | "FDS Test OK" logged | "FDS Test OK" | main.c | 504-509
BT-ST-020 | Self-Test | Self-test overall result (bitmask == 0) | CRITICAL | All sensors initialized and producing data | Parse "SELFTEST TERMINATED <value>" from log | test_result == 0 (all bits clear) | 0 (0b00000000) | mmt_selftest.c | 141
BT-ST-021 | Self-Test | Self-test TEMP_SENSOR_1 (b0) | MAJOR | Temperature Sensor 1 (LMT70 T1) initialized | Check bit 0 of selftest result bitmask | 1000 <= T1 <= 4100 (hundredths of degree C, i.e. 10.00C to 41.00C) | [1000, 4100] 1/100 degC | mmt_selftest.c
BT-ST-022 | Self-Test | Self-test TEMP_SENSOR_2 (b1) | MAJOR | Temperature Sensor 2 (LMT70 T2) initialized | Check bit 1 of selftest result bitmask | 1000 <= T2 <= 4100 (hundredths of degree C, i.e. 10.00C to 41.00C) | [1000, 4100] 1/100 degC | mmt_selftest.c
BT-ST-023 | Self-Test | Self-test EXT_FLASH (b2) | MAJOR | External Flash Memory (MX25R6435F / MT25QU256) initialized | Check bit 2 of selftest result bitmask | Reserved - not currently tested in selftest code | Reserved - not currently tested in selftest code | mmt_selftest.c
BT-ST-024 | Self-Test | Self-test PPG_INIT (b3) | MAJOR | PPG (FXI/Philips PSP) Initialization initialized | Check bit 3 of selftest result bitmask | mmt_fxi_process_is_initialized() returns true | mmt_fxi_process_is_initialized() returns true | mmt_selftest.c
BT-ST-025 | Self-Test | Self-test PPG_READ (b4) | MAJOR | PPG Data Reading (wearing detection measurement) initialized | Check bit 4 of selftest result bitmask | 50 <= PPG_value <= 65000 | [50, 65000] ADC counts | mmt_selftest.c
BT-ST-026 | Self-Test | Self-test ACC_STARTED (b5) | MAJOR | Accelerometer (LIS2DW12) Started initialized | Check bit 5 of selftest result bitmask | LIS2DW12_is_started() returns true | LIS2DW12_is_started() returns true | mmt_selftest.c
BT-ST-027 | Self-Test | Self-test ACC_RANGE (b6) | MAJOR | Accelerometer Data Range Validation initialized | Check bit 6 of selftest result bitmask | ¦AccX¦+¦AccY¦+¦AccZ¦ >= 50 AND NOT (¦AccX¦>20000 AND ¦AccY¦>20000 AND ¦AccZ¦>20000) | ¦AccX¦+¦AccY¦+¦AccZ¦ >= 50 AND NOT (¦AccX¦>20000 AND ¦AccY¦>20000 AND ¦AccZ¦>20000) | mmt_selftest.c
BT-ST-028 | Self-Test | Self-test PPG_TIMING (b7) | MAJOR | PPG/ACC Data Freshness (Timing Check) initialized | Check bit 7 of selftest result bitmask | time_now - last_measurement_time <= 180 seconds (3 minutes) | time_now - last_measurement_time <= 180 seconds (3 minutes) | mmt_selftest.c
BT-ALU-029 | ALU/FPU | ALU/FPU benchmark execution | MINOR | mmt_test_ALU_FPU() call uncommented in main.c | Check for ALU/FPU timing output in log | Timing data present for int32, float, double operations | 6 timing measurements expected | mmt_test_ALU_FPU.c | 206-210
BT-ALU-030 | ALU/FPU | ALU/FPU: Integer 32-bit addition/subtraction | MINOR | ALU/FPU benchmark running | Parse timing for int32_add_sub from log output | Execution time <= 50 ms (30000 iterations x 10 ops) | <= 50 ms | mmt_test_ALU_FPU.c
BT-ALU-031 | ALU/FPU | ALU/FPU: Integer 32-bit multiplication/division | MINOR | ALU/FPU benchmark running | Parse timing for int32_mul_div from log output | Execution time <= 150 ms (30000 iterations x 10 ops) | <= 150 ms | mmt_test_ALU_FPU.c
BT-ALU-032 | ALU/FPU | ALU/FPU: Single-precision float addition/subtraction | MINOR | ALU/FPU benchmark running | Parse timing for float_add_sub from log output | Execution time <= 50 ms (30000 iterations x 10 ops) | <= 50 ms | mmt_test_ALU_FPU.c
BT-ALU-033 | ALU/FPU | ALU/FPU: Single-precision float multiplication/division | MINOR | ALU/FPU benchmark running | Parse timing for float_mul_div from log output | Execution time <= 150 ms (30000 iterations x 10 ops) | <= 150 ms | mmt_test_ALU_FPU.c
BT-ALU-034 | ALU/FPU | ALU/FPU: Double-precision float addition/subtraction | MINOR | ALU/FPU benchmark running | Parse timing for double_add_sub from log output | Execution time <= 500 ms (30000 iterations x 10 ops) | <= 500 ms | mmt_test_ALU_FPU.c
BT-ALU-035 | ALU/FPU | ALU/FPU: Double-precision float multiplication/division | MINOR | ALU/FPU benchmark running | Parse timing for double_mul_div from log output | Execution time <= 1500 ms (30000 iterations x 10 ops) | <= 1500 ms | mmt_test_ALU_FPU.c
BT-TIM-036 | Boot Timing | Initialization phase duration | MAJOR | Boot started | Verify MMT_TASK_STOP_INITIALIZATION delay in task scheduler | STOP_INITIALIZATION scheduled at 5550 ms | 5550 ms | main.c | 775
BT-TIM-037 | Boot Timing | Time to sleep mode | MAJOR | Initialization phase completed | Verify MMT_TASK_SWITCH_TO_SLEEP delay in task scheduler | SWITCH_TO_SLEEP scheduled at 6050 ms | 6050 ms | main.c | 776
BT-TEMP-038 | Temperature Init | T1 temperature at boot within range | MAJOR | LMT70/TMP117 sensor initialized | Parse T1 value from boot log | 1000 <= T1 <= 4100 (10.00 C to 41.00 C) | [1000, 4100] (1/100 degC) | main.c | 728-736
BT-TEMP-039 | Temperature Init | T2 temperature at boot within range | MAJOR | LMT70 sensor initialized | Parse T2 value from selftest log | 1000 <= T2 <= 4100 (10.00 C to 41.00 C) | [1000, 4100] (1/100 degC) | mmt_selftest.c | 73-84
BT-HW-040 | Hardware Init | TMP117 I2C sensor detection | MINOR | I2C bus initialized | Parse "TMP sensor" value from log | TMP117 detected (value = 1) | TMP sensor 1 | main.c | 477
BT-INFO-041 | Firmware Info | Firmware version logged | INFO | FW_info module initialized | Parse FW version from "MMT 287-B2 STARTED FW_V" log line | FW version string present (format: Vx_y_z) | FW_Vx_y_z | main.c | 668
BT-INFO-042 | Firmware Info | Serial number logged | INFO | Serialization module initialized | Parse serial number from boot log | Serial number present (8 characters) | 8-char alphanumeric string | main.c | 666-668
BT-INFO-043 | Firmware Info | Reset reason captured | INFO | NRF_POWER module accessible | Parse reset reason from "reset reason X" log line | Reset reason value present (0=POR, 1=Pin, 2=WDT, 4=SW, 16=Lockup) | Integer value logged | main.c | 398, 669

## Boot Sequence
Order | Marker ID | Log Pattern | Mandatory | Phase | Description | Status
1 | BOOT_START | MMT application started at 0x40000 | YES | Power-On | MCU boots, reads reset reason from NRF_POWER->RESETREAS
2 | INIT_PHASE_1 | init 1. | YES | Core HW | Log, timers, power management initialized
3 | INIT_PHASE_2 | init 2. | YES | Watchdog/DFU | DCDC mode, watchdog, DFU bootloader initialized
4 | INIT_PHASE_3 | init 3. | YES | GPIOTE/FDS | GPIOTE driver initialized
5 | INIT_PHASE_4 | init 4. | YES | BLE | BLE stack (S140 SoftDevice) initialized
6 | INIT_PHASE_41 | init 41. | YES | State Machine | State machine, LED system initialized
7 | INIT_PHASE_42 | init 42. | YES | QSPI/I2C | QSPI flash (ext. memory) + I2C bus initialized
8 | INIT_PHASE_43 | init 43. | YES | I2C | I2C bus initialization confirmed
9 | INIT_PHASE_44 | init 44. | YES | SPI | SPI bus initialized (sensor channel)
10 | FDS_TEST_START | #### TEsting FDS: mmt_fds_test | YES | Flash Test | Flash Data Storage write/read test begins
11 | FDS_TEST_RESULT | FDS Test (OK¦NOK) | YES | Flash Test | FDS test result (OK = all 5 cycles pass)
12 | DFU_WAIT | waiting for potential DFU | YES | DFU Window | 1s DFU window: device waits for OTA firmware update
13 | INIT_PHASE_6 | init 6. | YES | Flash Log | External flash log system initialized
14 | APP_RUNNING | Running the app | YES | App Start | All core modules ready, entering application phase
15 | FW_VERSION | MMT 287-B2 STARTED FW_V | YES | FW Info | Firmware version and serial number logged
16 | MAIN_LOOP | MAIN LOOP STARTED | YES | Main Loop | Main event loop entered — boot complete

## Self-Test Bitmask
Bit | Mask (Hex) | Name | Sensor/Module | Pass Condition | Valid Range | Unit | Failure Impact | Notes
b0 | 0x01 | TEMP_SENSOR_1 | Temperature Sensor 1 (LMT70 T1) | 1000 <= T1 <= 4100 (hundredths of degree C, i.e. 10.00C to 41.00C) | 1000-4100 | 1/100 degC | Temperature monitoring unavailable for T1 channel
b1 | 0x02 | TEMP_SENSOR_2 | Temperature Sensor 2 (LMT70 T2) | 1000 <= T2 <= 4100 (hundredths of degree C, i.e. 10.00C to 41.00C) | 1000-4100 | 1/100 degC | Temperature monitoring unavailable for T2 channel
b2 | 0x04 | EXT_FLASH | External Flash Memory (MX25R6435F / MT25QU256) | Reserved - not currently tested in selftest code | N/A |  | External flash not verified (separate FDS test covers this) | Tested separately via mmt_fds_test()
b3 | 0x08 | PPG_INIT | PPG (FXI/Philips PSP) Initialization | mmt_fxi_process_is_initialized() returns true | N/A |  | Philips PSP/FXI optical processing not ready — no HR/SpO2
b4 | 0x10 | PPG_READ | PPG Data Reading (wearing detection measurement) | 50 <= PPG_value <= 65000 | 50-65000 | ADC counts | PPG sensor not producing valid data — wearing detection impaired
b5 | 0x20 | ACC_STARTED | Accelerometer (LIS2DW12) Started | LIS2DW12_is_started() returns true | N/A |  | Accelerometer not started — no motion data, step counting fails
b6 | 0x40 | ACC_RANGE | Accelerometer Data Range Validation | ¦AccX¦+¦AccY¦+¦AccZ¦ >= 50 AND NOT (¦AccX¦>20000 AND ¦AccY¦>20000 AND ¦AccZ¦>20000) | N/A |  | Accelerometer producing invalid data — stuck at zero or saturated | Detects both zero-output (sensor dead) and saturation (sensor stuck)
b7 | 0x80 | PPG_TIMING | PPG/ACC Data Freshness (Timing Check) | time_now - last_measurement_time <= 180 seconds (3 minutes) | N/A |  | Sensor data is stale (>3 min) — possible scheduling or timing issue

## FDS Test
Iteration | Write Value (i) | Write String | Expected Read (i) | Expected Read String | File ID | File Key | Pass Criteria
1 | 0 | MMT_FDS | 0 | MMT_FDS | 0x8123 | 0x7234 | d_read.i == d_write.i AND rc == NRF_SUCCESS
2 | 1 | MMT_FDS | 1 | MMT_FDS | 0x8123 | 0x7234 | d_read.i == d_write.i AND rc == NRF_SUCCESS
3 | 2 | MMT_FDS | 2 | MMT_FDS | 0x8123 | 0x7234 | d_read.i == d_write.i AND rc == NRF_SUCCESS
4 | 3 | MMT_FDS | 3 | MMT_FDS | 0x8123 | 0x7234 | d_read.i == d_write.i AND rc == NRF_SUCCESS
5 | 4 | MMT_FDS | 4 | MMT_FDS | 0x8123 | 0x7234 | d_read.i == d_write.i AND rc == NRF_SUCCESS
Overall Pass Criteria: | All 5 write/read cycles must return matching data with NRF_SUCCESS

## ALU-FPU Benchmark
Test | Data Type | Operations | Iterations | Ops/Iteration | Max Time (ms) | HW Support | Notes
int32_add_sub | int32 | ADD / SUB | 30000 | 10 | 50 | ALU (native) | Should be very fast, pure ALU operations
int32_mul_div | int32 | MUL / DIV | 30000 | 10 | 150 | ALU (native, HW mul) | Division is slower than multiplication on Cortex-M4
float_add_sub | float | FADD / FSUB | 30000 | 10 | 50 | FPU (FPv4-SP-D16) | Hardware FPU - should be nearly as fast as integer
float_mul_div | float | FMUL / FDIV | 30000 | 10 | 150 | FPU (FPv4-SP-D16) | FPU hardware support for single precision
double_add_sub | double | DADD / DSUB (emulated) | 30000 | 10 | 500 | Software emulation | NO hardware FPU for double - software emulation, significantly slower
double_mul_div | double | DMUL / DDIV (emulated) | 30000 | 10 | 1500 | Software emulation | Software emulation - slowest operation expected

## HW Init Checklist
# | Module | Init Function | Parameters / Device | Verification Check | Critical | Status
1 | Gpio | mmt_disable_all_gpio_pins |  | All GPIO pins set to default (low power) | NO
2 | Sleep Gpio | mmt_sleep_gpio_pins |  | Selective GPIO for charger + ACC | NO
3 | Power Mode | sd_power_mode_set | NRF_POWER_MODE_LOWPWR |  | NO
4 | Dcdc | sd_power_dcdc_mode_set | NRF_POWER_DCDC_ENABLE |  | NO
5 | Watchdog | System_task_initWatchdog |  |  | YES
6 | Timers | timers_init |  | app_timer_init returns NRF_SUCCESS | NO
7 | Power Mgmt | power_management_init |  | nrf_pwr_mgmt_init returns NRF_SUCCESS | NO
8 | Gpiote | nrf_drv_gpiote_init |  |  | NO
9 | Fds | mmt_fds_init |  |  | YES
10 | Ble | mmt_ble_all_init |  |  | YES
11 | Qspi | qspi_init | External Flash SPI channel 2 |  | YES
12 | I2C | i2c_init |  | mmt_i2c_master_init returns NRF_SUCCESS | YES
13 | Spi | spi_init | Sensor SPI channel 1 |  | YES
14 | Tmp117 | mmt_TMP117_init |  | Returns detection status boolean | NO
15 | Fpu Irq | NVIC_SetPriority | FPU_IRQn, APP_IRQ_PRIORITY_HIGH |  | YES
16 | Clock | mmt_clock_init | mmt_second_event_handler |  | YES
17 | Battery | mmt_battery_mgmt_Init |  |  | NO
18 | Wearing | mmt_wearing_init |  |  | NO
19 | Acc Interrupt | ACC_init | MMT_ACC_DATA_READY_PIN |  | NO
20 | Flash Log | mmt_log_init |  |  | NO
21 | State Machine | mmt_state_machine_init |  |  | NO
22 | Leds | mmt_LEDS_init |  |  | NO
23 | Greenteg | mmt_GreenTeg_init |  |  | NO
24 | Nibp | mmt_nibp_init |  |  | NO
25 | Cardiac Arrest | corsano_cardiac_arrest_init |  |  | NO
26 | Idle Detection | mmt_idle_detection_init |  |  | NO

## Task Scheduler
Task ID | Task Name | Delay (ms) | Delay (ticks @32.768kHz) | Description | Execution Order | Critical Path
50 | MMT_TASK_RUN_ACC | 10 | 327 | Start accelerometer | 1 | NO
95 | MMT_TASK_SET_HRM_MODE | 50 | 1638 | Start heart rate monitoring | 2 | NO
128 | MMT_TASK_BIOZ_START_CALIBRATION_1 | 70 | 2293 | Start BioZ calibration | 3 | NO
139 | MMT_TASK_STOP_INITIALIZATION | 5550 | 181862 | End initialization phase | 4 | YES
80 | MMT_TASK_SWITCH_TO_SLEEP | 6050 | 198246 | Switch to sleep mode | 5 | YES